## Slide 1: SFMR operations on the WC130J during the 2009 Hurricane Season

## Slide 2: 2009 SFMR system performance
Two issues affected system performance/accuracy during the 2009 Season
  Water intrusion into the radome area (major issue)
    Fixed Sep 09
  Lack of working interference rejection algorithm (minor issue)
    Fixed late Nov 09

## Slide 3: 2009 SFMR system performance
Water intrusion
  Anomaly first noticed on 0408E Felicia mission
  Affected nearly 50% of missions flown during season
  Cause was retrofit of new SFMR antenna
  Fix was fairly simple
  Determining what was happening was not simple

## Slide 4: 2009 SFMR system performance
Old antenna
New Antenna

## Slide 5: 2009 SFMR system performance
Drain Holes

## Slide 6: 2009 SFMR system performance
In heavy rain situations, water would enter the rear drain holes and flow on the inside top of the antenna

## Slide 7: 2009 SFMR system performance
Chart: Hurricane Jimena 1 Sep 09
Example of water leakage into radome area

## Slide 8: 2009 SFMR system performance
Chart: Hurricane Jimena, 1 Sep 09 US010
Example of water leakage into radome area

## Slide 9: 2009 SFMR system performance
Fix was to seal the antenna drain holes
All instruments sealed late Sep 09
Excellent results in Hurricane IDA (Nov)

## Slide 10: 2009 SFMR system performance
Chart: 0311A IDA US010
Example of no water intrusion

## Slide 11: 2009 SFMR system performance
Chart: 0111A IDA US012
Example of no water intrusion

## Slide 12: 2009 SFMR system performance
70% of the missions with a large rain profile had water intrusion problems
  50% of Bill missions, Claudette, Jimena and one Rick mission
After occurrence, calculated winds were too high
  Data inconsistent, led to frustrated operators and forecasters
Calculated winds up to first fix were nearly always correct. Calculated winds after first fix were generally good in the eye wall/high wind/rain but bad in light wind areas.
In some cases it did not occur. (Don’t know why)
Only clue operators had was indication of higher RMS error and non verifying visual estimates. (Sometimes impossible on night missions)
Antenna areas sealed late Sep 09. No occurrences since that time.

## Slide 13: 2009 SFMR system performance
Other enhancements for 2010 season
  Processor Software update V1.15
    Installed Nov 09
    Includes working rejection algorithm
  Operator Software visual application enhancement Nov 09
    Enhanced visual que’s to equipment operation
    Enhanced data download utility
ProSensing/AFRC/NOAA AOC working to reduce differences in Platform SFMR calculations WC-103J VS P-3